## Slide 1: Qingping Chen Division of Medical Physics, Dept. Of Radiology, University Medical Center Freiburg, Germany March 25, 2024
Basic Sequence Exercises
Notes: Change your last name. some hint there. Introduction and wrap up. Build your own sequences. Opportunity to test on scanner.

## Slide 2: Prerequisites
Programming tools
  Matlab software needs to be installed in your computer.
  Basic familiarity with Matlab programming is required.
Pulseq software
  Please download the Pulseq software from https://github.com/pulseq/pulseq. Install Pulseq software in Matlab by adding its directory and subdirectories to Matlab's path.
Text compare tool
  We recommend using a text comparison tool to compare the sequences within the subsequent steps to visualize the changes that occur at each step. Meld software can be used for text comparison and can be downloaded from https://meldmerge.org/.
Basic sequence exercises
  Basic MR spectroscopy: from FID to 1D spin-echo and 1D gradient-echo
  Basic MR imaging: from FID to basic GRE to FLASH

## Slide 3: ex01_fid
ADC
Gx
Gy
Gz
RF
1
2
RF
ex
3
delayTE
TE
delayTR
TR
4
kx
ky
0
0
T2*-w FID
Task: ex01_fid is the simplest FID sequence. Run it and get familiar with Pulseq programming.

## Slide 4: ex02_fid2se
ADC
Gx
Gy
Gz
RF
1
2
RF
ex
3
delayTE
TE
delayTR
TR
4
T2*-w FID
T2*: macroscopic and microscopic field inhomogeneity
T2: microscopic field inhomogeneity
T2 > T2*

## Slide 5: ex02_fid2se
ADC
Gx
Gy
Gz
RF
1
2
RF
ex
RF
ref
3
4
delayTE1
TE/2
TE/2
delayTR
TR
5
kx
ky
0
0
T2-w
delayTE2
Task: insert a 180-deg RF pulse between RF_ex and ADC to ex01_fid to reduce macroscopic inhomogeneity effect. The acquired signal is the so-called spin-echo.

## Slide 6: ex03_se_crushers
ADC
Gx
Gy
Gz
RF
1
2
RF
ex
RF
ref
3
4
delayTE1
TE/2
TE/2
delayTR
TR
5
T2-w
180° pulse is typically not perfect. Crushers to suppress unwanted FID
Unwanted FID
delayTE2

## Slide 7: ex03_se_crushers
ADC
Gx
Gy
Gz
RF
1
2
RF
ex
RF
ref
3
4
delayTE1
TE/2
TE/2
delayTR
TR
crushers
5
kx
ky
0
0
T2-w
delayTE2
Task: insert a pair of crushers before and after RF_ref in ex02_fid2se to dephase unwanted signal.

## Slide 8: ex04_fid2gre1d
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
0
-3
-2
-1
1
2
Task: insert a readout gradient and its pre-dephaser in ex01_fid to generate a gradient recalled echo (GRE).

## Slide 9: ex11_gre1d2gre2d
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Task: insert a phase-encoding gradient in ex04_fid2gre1d for 2D spatial encoding. The extended sequence is the most basic 2D GRE sequence.

## Slide 10: ex12_gre2d_lbl
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Label for PE
Task: Run ex12_gre2d_lbl to get familiar with ADC labeling for ICE online reconstruction in the afternoon session.
Set PE label counter:
mr.makeLabel('SET', 'LIN', PElbl(i))
Increment PE label counter by 1:
mr.makeLabel(‘INC', 'LIN', 1)

## Slide 11: ex13_gre2d_gradSpoil
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Label for PE
spoiler
spoiler
Task: Add gradient spoiling in slice-selective and readout directions in ex12_gre2d_lbl to dephase residual tranverse magnetization.

## Slide 12: ex14_gre2d_PErefocus
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Label for PE
spoiler
spoiler
Task: Insert a phase-encoding backblip in ex13_gre2d_gradSpoil, such that the end point of all k-space trajectory is the same to further suppress artifacts.

## Slide 13: ex15_gre2d_RFspoil
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Label for PE
spoiler
spoiler
Task: another artifact can arise from the fact that we always flip the magnetization in the same direction in each repetition. insert an altered RF/ADC phase in each repetition in ex14_gre2d_PErefocus.

## Slide 14: ex16_flash_fasterTiming (optional)
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
delayTR
TR
4
5
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
Label for PE
spoiler
spoiler
Task: With GRE combined with gradient spoiling and RF spoiling, we have generated the famous FLASH sequence. We can make this sequence faster by doing a "gradient surgery“ for ex15_gre2d_RFspoil. Could you try to do this surgery by yourself?

## Slide 15: ex16_flash_fasterTiming (optional)
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
TR
4
spoiler
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
spoiler
Gradient “surgery”:
mr.splitGradientAt(…)
mr.makeExtendedTrapezoidArea(…)
mr.addGradients(…)
Label for PE

## Slide 16: ex17_flash_accelerateComputation (optional)
ADC
Gx
Gy
Gz
RF
1
2
3
delayTE
TE
TR
4
spoiler
kx
ky
-3
-2
-1
0
2
1
0
-3
-2
-1
1
2
spoiler
Label for PE
To accelerate seq.addBlock(…):
seq.registerGradEvent(…)
seq.registerRfEvent(…)

## Slide 17
Caution! Possible source of errors!
After the object is registered, the seq.addBlock(…) will never search the library for consistency.
RF pulse with changing phase for RF spoiling
[~, rf.shapeIDs] = seq.registerRfEvent(rf) ;
rf.id = seq.registerRfEvent(rf) ; % NO GO EXAMPLE!!!
ex17_flash_accelerateComputation (optional)